SOP: Automated Algae Counting in ImageJ

Date: November 21, 2025

Scope: Fluorescence Microscopy & Hemocytometer Analysis

1. Purpose

To automate the counting of algal cells (e.g., Symbiodiniaceae, Phytoplankton) from fluorescence microscope images. This protocol uses a custom ImageJ macro to:

- Standardize the Count Area: Automatically crops the image to the exact volume of the Hemocytometer Central Square (0.1 µL).

- Remove Bias: Uses mathematical thresholding rather than human eye estimation.

- Handle Variable Replicates: Automatically adjusts to any number of photos per sample.

2. The Logic: How the Parameters Were Calculated

Before running the script, it is important to understand why the script behaves the way it does.

A. The "Cut" (Area of Interest)

To calculate concentration (Cells/mL), we must know the exact volume being counted.

- Standard: The Central Square of a Neubauer Hemocytometer is 1mm x 1mm.

- Calibration: We measured the Central Square on the 10X brightfield image (1535 px wide).

- Resolution Correction: The fluorescence images were taken at a lower resolution (Binning 3).

  - 1535 \{ (original grid)} / 3 = \mathbf{512 \text{ pixels}}$.

  - Result: The script automatically creates a 512x512 pixel box. Anything inside this box represents exactly 0.1 µL of sample volume.

B. The Threshold (Detecting Cells)

- Method: "Triangle".

- Why: Our samples contain "Dim" cells (faint fluorescence). The standard method ("Otsu") was too strict. "Triangle" effectively detects low-intensity signals without picking up background static.

C. The Size Filter (Ignoring Dust)

- Setting: > 50 pixels.

- Why: Analysis showed that background dust is approx. 14 pixels. Setting the filter to 50 ensures dust is ignored, while even the smallest algae (approx. 240px) are preserved.

3. Instructions

Phase 1: Preparation

- Organize Files: Place all your .tif, .jpg, or .png images into a single Input Folder.

  - Naming Convention: SampleID_Replicate.tif (e.g., 401_1.tif, 401_2.tif). The script groups replicates automatically.

- Create Output Folder: Create an empty folder named "Results" or "Counted_Images".

Phase 2: Running the Script

- Open Fiji / ImageJ.

- Drag the script file (AlgaeCounter.ijm) into the Fiji bar (or go to Plugins > Macros > Run).

- Select Input Folder: Choose the folder with your images.

- Select Output Folder: Choose your results folder.

Phase 3: The "Eraser" Step (Interactive)

The script will open the images one by one. For EACH image:

- The script will pause and display a prompt.

- Look for the Scale Bar (or any large debris clumps).

- Use the mouse to Draw a Box over the text/scale bar.

- Click OK on the prompt window.

  - The script will verify: It paints the box black (ignoring it), counts the algae, and moves to the next image.

Phase 4: Results & Quality Control

- Extracting Data to Excel (UPDATED LAYOUT):

  - When the script finishes, a window named "Log" will appear with your data table.

  - Click anywhere inside the "Log" window to activate it.

  - Press Ctrl + A (Select All) and Ctrl + C (Copy).

  - Paste into Excel.

  - NOTE: To support variable numbers of photos, the columns are now ordered:

Sample Name -> Concentration -> Average -> Count 1 -> Count 2 -> Count 3...

(Column B will always be your Final Result).

- Visual Verification:

  - Go to your Output Folder.

  - Open the images named Checked_....

  - Cyan Numbers: These are the cells the computer counted.

  - Verify that the scale bar is gone and faint cells are numbered.

4. The Script Code

Copy the code below and save it as AlgaeCounter.ijm.

/* * SOP SCRIPT: AUTOMATED ALGAE COUNTER (Variable Replicates) * Lab Protocol: 11/2025 * * PARAMETERS: * - Crop Size: 512x512 (Calculated for 10X Binning 3) * - Min Size: 50 pixels (Filters dust <14px) * - Threshold: Triangle (Optimized for dim fluorescence) */ macro "SOP Algae Variable [F2]" { // --- CONFIGURATION --- minSize = 50; boxSize = 512; threshMethod = "Triangle"; // --- SETUP --- run("Close All"); print("\\Clear"); roiManager("Reset"); // Folder Selection inputDir = getDirectory("Select the INPUT Folder (Images)"); outputDir = getDirectory("Select the OUTPUT Folder (Results)"); list = getFileList(inputDir); Array.sort(list); // Excel Header print("Sample No\tConc (Cells/mL)\tAverage\tCount 1\tCount 2\tCount 3\t(etc...)"); currentSample = ""; counts = newArray(0); // Interactive Mode (BatchMode OFF to allow erasing) setBatchMode(false); for (i = 0; i < list.length; i++) { filename = list[i]; if (endsWith(filename, ".tif") ¦¦ endsWith(filename, ".jpg") ¦¦ endsWith(filename, ".png")) { // 1. Parse Sample ID dotIndex = lastIndexOf(filename, "."); nameNoExt = substring(filename, 0, dotIndex); parts = split(nameNoExt, "_"); sampleID = parts[0]; // 2. Data Grouping if (sampleID != currentSample) { if (currentSample != "") { printRow(currentSample, counts); } currentSample = sampleID; counts = newArray(0); } // 3. Image Prep open(inputDir + filename); originalTitle = getTitle(); // Auto-Crop to 0.1 uL Volume makeRectangle(getWidth()/2 - boxSize/2, getHeight()/2 - boxSize/2, boxSize, boxSize); run("Crop"); // 4. USER INTERACTION: Scale Bar Eraser setTool("rectangle"); waitForUser("SOP Action", "Draw a box over the SCALE BAR text.\nThen click OK."); if (selectionType() != -1) { run("Set...", "value=0"); // Paint black run("Select None"); } // 5. Analysis run("Duplicate...", "title=Detection_Mask"); run("8-bit"); run("Subtract Background...", "rolling=50"); setAutoThreshold(threshMethod + " dark"); run("Convert to Mask"); run("Watershed"); // Count roiManager("Reset"); run("Analyze Particles...", "size=" + minSize + "-Infinity circularity=0.30-1.00 exclude add"); thisCount = roiManager("count"); counts = Array.concat(counts, thisCount); // 6. Save Evidence selectWindow(originalTitle); run("8-bit"); roiManager("Show All with labels"); run("Labels...", "color=Cyan font=14 show use draw"); run("Flatten"); saveAs("Jpeg", outputDir + "Checked_" + filename); close("*"); roiManager("Reset"); } } // Print Final Row if (currentSample != "") { printRow(currentSample, counts); } showMessage("SOP Complete", "All images processed.\n\n1. Click inside the LOG window.\n2. Press Ctrl+A (Select All).\n3. Press Ctrl+C (Copy).\n4. Paste into Excel."); } function printRow(sample, cArr) { sum = 0; n = cArr.length; // Calculate Stats for (k=0; k<n; k++) { sum = sum + cArr[k]; } avg = 0; if (n > 0) { avg = sum / n; } conc = avg * 10000; // Stats First Layout line = sample + "\t" + conc + "\t" + avg; // Append Variable Counts for (k=0; k<n; k++) { line = line + "\t" + cArr[k]; } print(line); }